## Cover Page
Functional Equivalents
ECRI has researched the SARS-CoV-2 Test Kits to compare feature functionality and availability.
Last Updated: | 2020-04-21
For additional information contact: | priceguide@ecri.org

## Tests
| Manufacturer | Vendor | Model | Catalog # | Description | FDA-Approval | Website | IFU | Sample Type | Analysis Technique | Point-of-Care? | Platform/Instrument Compatibility | Advertised Result Time | CLIA Complexity | Manufacturer Listed Compatible Transport Media | Number of Tests/Reactions
Tests with FDA EUA | Abbott Molecular | Abbott | Abbott RealTime SARS-COV-2 assay | 09N77-095 | The Abbott RealTime SARS-CoV-2 assay is a real-time (rt) reverse transcriptase (RT) polymerase chain reaction (PCR) test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in nasopharyngeal (NP) and oropharyngeal (OP) swabs from patients suspected of COVID-19 by their health care provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests. | EUA | https://www.molecular.abbott/us/en/products/infectious-disease/RealTime-SARS-CoV-2-Assay | https://www.molecular.abbott/sal/9N77-095_SARS-CoV-2_US_EUA_Amp_PI.pdf | Swab | PCR | No | Abbott m2000 System | Not specified | High | Not specified | 96
| Abbott Rapid Diagnostics | Abbott | ID NOW COVID-19 assay | 190-000 | ID NOW COVID-19 assay performed on the ID NOW Instrument is a rapid molecular in vitro diagnostic test utilizing an isothermal nucleic acid amplification technology intended for the qualitative detection of nucleic acid from the SARS-CoV-2 viral RNA in direct nasal, nasopharyngeal or throat swabs and nasal, nasopharyngeal or throat swabs eluted in viral transport media from individuals who are suspected of COVID-19 by their healthcare provider. Testing is authorized for laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform moderate complexity/high complexity tests. The ID NOW COVID-19 assay is also authorized to be distributed and used in patient care settings outside of the clinical laboratory environment. | EUA | https://www.alere.com/en/home/product-details/id-now-covid-19.html | https://ensur.invmed.com/ensur/contentAction.aspx?key=ensur.514295.S2R4E4A3.20200328.10292.4358420 | Swab | NAA | Yes | ID NOW | 13 | Not specified | Amie’s Media Dulbecco’s Modified Eagles’ Medium (D-MEM) Hank’s Balanced Salt Solution M4 Media M4-RT Media M5 Media M6 Media Phosphate Buffered Saline Saline Stuart’s Media Universal Transport Media Starplex Multitrans | 24
| Atila Biosystems, Inc | Atila Biosystems, Inc | iAMP COVID-19 Detection Kit | iAMP-COVID19-100 | The iAMP COVID-19 Detection Kit is a real-time fluorescent reverse transcription isothermal assay intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in nasal, nasopharyngeal (NP), and oropharyngeal (OP) swabs from patients suspected of COVID-19 by their health care provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests. | EUA | https://atilabiosystems.com/covid-19-test-2/ | https://www.fda.gov/media/136870/download | Swab | PCR | No | Atila Power-Gene 9600 Plus, Bio-Rad CFX96, ABI7500, | 60 | High | Not specified | 100
| Avellino Lab USA, Inc | Avellino Lab USA, Inc. | AvellinoCoV2 test | Not specified | The AvellinoCoV2 test is a real-time RT-PCR test intended for the qualitative detection of nucleic acid from SARS-CoV-2 in nasopharyngeal and oropharyngeal swab specimens collected from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to Avellino Lab USA, Inc., that is certified under Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a to perform high complexity tests. | EUA | https://avellinocoronatest.com/product/ | Not specified | Swab | PCR | Not specified | Applied Biosystems 7500 | 24-48 hours (turnaround) | High | Avellino Transport Media | 1
| Becton Dickinson | Becton Dickinson | BioGX SARS-CoV-2 Reagents for BD MAX System | 444213 | The BioGX SARS-CoV-2 Reagents for BD MAX™ System is a real-time RT-PCR test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in nasopharyngeal and oropharyngeal swab samples from individuals suspected of COVID-19 by their healthcare provider. | EUA | https://www.bd.com/en-us/offerings/capabilities/molecular-diagnostics/molecular-tests/biogx-sars-cov-2-reagents | https://www.fda.gov/media/136816/download | Swab | PCR | No | BD MAX | Not specified | Moderate, High | BD UVT, Copan UTM | 24
| BGI Genomics Co. Ltd | BGI Genomics Co. Ltd | Real-Time Fluorescent RT-PCR Kit for Detecting SARS-2019-nCoV | HW5105 | The Real-Time Fluorescent RT-PCR Kit for Detecting SARS-2019-nCoV is an in vitro diagnostic real-time reverse transcription-PCR assay for the qualitative detection of SARS-CoV-2 nucleic acids in throat swabs and bronchoalveolar lavage fluid (BALF) from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 USC §263a, to perform high complexity tests. | EUA | https://www.bgi.com/us/2019-ncov-real-time-fluorescent-rt-pcr-kit-ivd/ | https://www.bgi.com/us/wp-content/uploads/sites/2/2020/03/EUA-Real-Time-Fluorescent-SARS-2019-nCoV-BGI-IFU.pdf | Swab | PCR | No | Applied Biosystems 7500 | 180 | High | Not specified | 50
| BioFire Defense, LLC | BioFire Defense, LLC | BioFire COVID-19 Test Kit | 423744 | The BioFire COVID-19 Test is a nested multiplexed real-time RT-PCR test intended for the qualitative detection of nucleic acid from SARS-CoV-2 in nasopharyngeal swabs in transport media from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to U.S. laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform moderate complexity tests, and in U.S. laboratories certified under CLIA to perform high complexity tests, or in similarly qualified non-U.S. laboratories. | EUA | https://www.biofiredefense.com/covid-19test/ | http://www.biofiredefense.com/wp-content/uploads/2020/04/BioFire-COVID-19-Instructions-for-Use-DFA2-PRT-0080.pdf | Swab | PCR | No | FilmArray 2.0 and FilmArray Torch | 60 | Moderate, High | Not specified | 30
| BioFire Defense, LLC | BioFire Defense, LLC | BioFire COVID-19 Test Kit | 423745 | The BioFire COVID-19 Test is a nested multiplexed real-time RT-PCR test intended for the qualitative detection of nucleic acid from SARS-CoV-2 in nasopharyngeal swabs in transport media from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to U.S. laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform moderate complexity tests, and in U.S. laboratories certified under CLIA to perform high complexity tests, or in similarly qualified non-U.S. laboratories. | EUA | https://www.biofiredefense.com/covid-19test/ | http://www.biofiredefense.com/wp-content/uploads/2020/04/BioFire-COVID-19-Instructions-for-Use-DFA2-PRT-0080.pdf | Swab | PCR | No | FilmArray 2.0 and FilmArray Torch | 60 | Moderate, High | Not specified | 6
| Centers for Disease Control and Prevention | Centers for Disease Control and Prevention | CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel | Not specified | Not specified | EUA | Not specified | Not specified | Swab | PCR | No | Not specified | Not specified | Not specified | Not specified | Not specified
| Cepheid | Cepheid | Xpert Xpress SARS-CoV-2 test | Not specified | Cepheid has developed an automated molecular test for the qualitative detection of SARS-CoV-2, the virus that causes COVID-19. The test leverages the design principles of our current Xpert® Xpress Flu/RSV cartridge technology, in which multiple regions of the viral genome are targeted | EUA | https://www.cepheid.com/Package%20Insert%20Files/Xpress-SARS-CoV-2-PI/302-3750-rev-A-PACKAGE-INSERT-EUA-XPRESS-SARS-COV2.pdf | https://www.cepheid.com/Package%20Insert%20Files/Xpress-SARS-CoV-2-PI/302-3750-rev-A-PACKAGE-INSERT-EUA-XPRESS-SARS-COV2.pdf | Swab | PCR | Not specified | GeneXpert Xpress System, or GeneXpert Dx, or GeneXpert Infinity System | 45 | Waived (GeneXpert Express System), Moderate, High | Not specified | 10
| Co-Diagnostics, Inc | Co-Diagnostics, Inc | Logix Smart Coronavirus Disease 2019 (COVID-19) Kit | Not specified | The Logix Smart Coronavirus Disease 2019 (COVID-19) test is a real-time RT-PCR test using a proprietary technology called CoPrimers™ (Satterfield, 2014)(Poritz & Ririe, 2014) intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in lower respiratory tract fluids (e.g. bronchoalveolar lavage, sputum, tracheal aspirate), and upper respiratory tract fluids (e.g. nasopharyngeal and oropharyngeal swabs) from individuals with signs and symptoms of infection who are suspected of having COVID-19. The test is available to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests, or by similarly qualified non-U.S. laboratories. | EUA | http://codiagnostics.com/products/diagnostic-solutions/logix-smart-covid19/ | http://codiagnostics.com/wp-content/uploads/2020/04/Instructions-for-Use-Logix-Smart™-COVID-19-EUA-CE-IVD-Apr-2020.pdf | Swab | PCR | No | Not specified | 90 | High | Not specified | Not specified
| DiaCarta, Inc | DiaCarta, Inc | QuantiVirus SARS-CoV-2 Test kit | DC-11-0007 | The QuantiVirus™ SARS-CoV-2 Test kit is a real-time RT-PCR test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in nasal swabs, nasopharyngeal swabs, oropharyngeal swabs, and sputum from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests. | EUA | https://diacarta.com/products/coronavirus-test | https://diacarta.com/get-download-access/download-id/16356 | Swab | PCR | No | Thermo Fisher QuantStudio 5, Applied Biosystems 7500 Fast Dx, Bio-Rad CFX 384 | 60 | High | Not specified | 24
| DiaCarta, Inc | DiaCarta, Inc | QuantiVirus SARS-CoV-2 Test kit | DC-11-0008 | The QuantiVirus™ SARS-CoV-2 Test kit is a real-time RT-PCR test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in nasal swabs, nasopharyngeal swabs, oropharyngeal swabs, and sputum from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests. | EUA | https://diacarta.com/products/coronavirus-test | https://diacarta.com/get-download-access/download-id/16356 | Swab | PCR | No | Thermo Fisher QuantStudio 5, Applied Biosystems 7500 Fast Dx, Bio-Rad CFX 384 | 60 | High | Not specified | 48
| DiaCarta, Inc | DiaCarta, Inc | QuantiVirus SARS-CoV-2 Test kit | DC-11-0009 | The QuantiVirus™ SARS-CoV-2 Test kit is a real-time RT-PCR test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in nasal swabs, nasopharyngeal swabs, oropharyngeal swabs, and sputum from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests. | EUA | https://diacarta.com/products/coronavirus-test | https://diacarta.com/get-download-access/download-id/16356 | Swab | PCR | No | Thermo Fisher QuantStudio 5, Applied Biosystems 7500 Fast Dx, Bio-Rad CFX 384 | 60 | High | Not specified | 480
| DiaSorin Molecular LLC | DiaSorin Inc | Simplexa COVID-19 Direct assay | MOL4150 | The DiaSorin Molecular Simplexa COVID-19 Direct real-time RT-PCR assay is intended for use on the LIAISON MDX instrument for the in vitro qualitative detection of nucleic acid from severe acute respiratory syndrome coronavirus 2 (SARS-CoV-2) in nasopharyngeal swabs (NPS) from individuals suspected of COVID-19 by their healthcare provider. The Simplexa COVID-19 Direct assay is an aid in the diagnosis of SARS-CoV-2 infection. | EUA | https://molecular.diasorin.com/us/kit/simplexa-covid-19-direct-kit/ | https://www.fda.gov/media/136286/download | Swab | PCR | No | LIAISON MDX | 60 | Moderate, High | Copan Universal Transport Media, BD Universal Viral Transport, Remel M5, Remel M6, Copan Eswab, Puritan UniTranz-RT, Saline | 24
| Fosun Pharma USA Inc. | Fosun Pharma USA Inc. | Fosun COVID-19 RT-PCR Detection Kit | Not specified | Not specified | EUA | Not specified | Not specified | Swab | PCR | Not specified | Not specified | Not specified | Not specified | Not specified | Not specified
| GenMark Diagnostics, Inc. | GenMark Diagnostics, Inc. | ePlex SARS-CoV-2 Test | EA008212 | The ePlex SARS-CoV-2 Test is intended for the qualitative detection of SARS-CoV-2 nucleic acid in nasopharyngeal swab specimens (NPS) eluted in viral transport media (VTM) collected from individuals suspected of COVID-19 by their health care provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. § 263a, to perform moderate and high complexity tests. | EUA | https://www.genmarkdx.com/solutions/panels/eplex-panels/eplex-sars-cov-2-test/ | https://www.fda.gov/media/136282/download | Swab | PCR | No | GenMark ePlex | 2 | Moderate, High | Not specified | 12
| GenoSensor, LLC | GenoSensor, LLC | GS™ COVID-19 RT-PCR KIT | 2702-22 | GSTM COVID-19 RT-PCR KIT is a real-time reverse transcription polymerase chain reaction (rRT-PCR) test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in nasopharyngeal/oropharyngeal swabs, nasal swabs and mid-turbinate swabs from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests. | EUA | http://genosensorcorp.com/ | http://genosensorcorp.com/COVIDpdf/GenoSensor_IFU.pdf | Swab | PCR | No | Applied Biosystems 7500 | Not specified | High | Not specified | 22
| GenoSensor, LLC | GenoSensor, LLC | GS™ COVID-19 RT-PCR KIT | 2702-94 | GSTM COVID-19 RT-PCR KIT is a real-time reverse transcription polymerase chain reaction (rRT-PCR) test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in nasopharyngeal/oropharyngeal swabs, nasal swabs and mid-turbinate swabs from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests. | EUA | http://genosensorcorp.com/ | http://genosensorcorp.com/COVIDpdf/GenoSensor_IFU.pdf | Swab | PCR | No | Applied Biosystems 7500 | Not specified | High | Not specified | 94
| Gnomegen LLC | Gnomegen LLC | Gnomegen COVID-19 RT-Digital PCR Detection Kit | CV0202 | The Gnomegen COVID-19 RT-Digital PCR Detection kit is a real-time digital PCR in vitro diagnostic test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 virus in nasal, nasopharyngeal, and oropharyngeal swab specimens collected from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to laboratories - certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests. | EUA | https://gnomegendx.com/ | https://www.fda.gov/media/136738/download | Swab | PCR | No | QuantStudio 3D digital PCR System | 90 | High | Not specified | 48
| Hologic Inc | Hologic Inc | Panther Fusion SARS-CoV-2 test | PRD-04303 | The Panther Fusion SARS-CoV-2 Assay is a real-time RT-PCR in vitro diagnostic test intended for the qualitative detection of RNA from SARS-CoV-2 isolated and purified from nasopharyngeal (NP) and oropharyngeal (OP) swab specimens obtained from individuals who meet COVID-19 clinical and/or epidemiological criteria. The Panther Fusion SARS-CoV-2 Assay is for use only under Emergency Use Authorization (EUA) in the US laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C.§263a, to perform high complexity tests. | EUA | https://www2.hologic.com/pantherfusion-sars-cov-2 | https://www.fda.gov/media/136156/download | Swab | PCR | No | Panther Fusion System | Not specified | High | Remel MicroTest M4, M4RT, M5, M6, Copan Universal Transport Medium, BD Universal Transport Medium | 96
| InBios International, Inc | InBios International, Inc | Smart Detect SARS-CoV-2 rRT-PCR Kit | COV2-E | Smart DetectTM SARS-CoV-2 rRT-PCR Kit is a real-time RT-PCR test intended for the qualitative detection of nucleic acid from severe acute respiratory syndrome-related coronavirus 2 (SARS-CoV-2) in human nasopharyngeal swab, anterior nasal swab and mid-turbinate nasal swab specimens from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests. | EUA | https://inbios.com/smart-detecttm-sars-cov-2-rrt-pcr-kit/ | http://i8sit3w4v3z1h99oi1gmrx61-wpengine.netdna-ssl.com/wp-content/uploads/2020/04/LBL-0109-00-900251-EUA-Smart-Detect-SARS-CoV-2-rRT-PCR-Kit-Package-Insert.pdf | Swab | PCR | No | Applied Biosystems 7500, Bio-Rad CFX96 | 240 | High | Not specified | 48
| Ipsum Diagnostics, LLC | Ipsum Diagnostics, LLC | COV-19 IDx assay | Not specified | Not specified | EUA | https://www.ipsumdiagnostics.com/for-physicians/supply-order-form-covid-19/ | Not specified | Swab | PCR | Not specified | Not specified | Not specified | High | Not specified | Not specified
| KorvaLabs Inc. | KorvaLabs Inc. | Curative-Korva SARS-Cov-2 Assay | Not specified | Not specified | EUA | Not specified | Not specified | Swab | PCR | Not specified | Not specified | Not specified | High | Not specified | Not specified
| Laboratory Corporation of America (LabCorp) | Laboratory Corporation of America (LabCorp) | COVID-19 RT-PCR Test | Not specified | Not specified | EUA | https://www.labcorp.com/tests/139900/2019-novel-coronavirus-covid-19-naa | Not specified | Swab | PCR | No | Not specified | 24-48 hours (turnaround) | High | Not specified; exclusions are provided | Not specified
| Luminex Molecular Diagnostics Inc | Luminex Molecular Diagnostics Inc | NxTAG CoV Extended Panel Assay | I051C0447 | The NxTAG CoV Extended Panel (NxTAG CoV) is a qualitative, multiplex, nucleic acid test for the detection of SARS-CoV-2 in nasopharyngeal swab samples. Designed for use on the MAGPIX System, the NxTAG CoV panel detects three separate gene targets to ensure accurate, reliable results. | EUA | https://www.luminexcorp.com/nxtag-respiratory-pathogen-panel/ | http://info.luminexcorp.com/covid19 | Swab | PCR | No | Luminex MAGPIX | Not specified | High | Not specified | 96
| Maccura Biotechnology (USA) LLC | Maccura Biotechnology (USA) LLC | SARS-CoV-2 Fluorescent PCR Kit | BUSGN7101109 | SARS-CoV-2 Fluorescent PCR Kit is a real-time RT-PCR test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in orpharyngeal swab and sputum specimens from individuals with signs and symptoms of infection who meet CDC criteria for SARS-CoV-2 testing. | EUA | https://www.maccura.com/en/ | https://www.maccura.com/resources/98099844450616688.pdf | Swab | PCR | No | Applied Biosystems 7500 | 120 | High | Not specified | 32
| Maccura Biotechnology (USA) LLC | Maccura Biotechnology (USA) LLC | SARS-CoV-2 Fluorescent PCR Kit | BUSGN7102109 | SARS-CoV-2 Fluorescent PCR Kit is a real-time RT-PCR test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in orpharyngeal swab and sputum specimens from individuals with signs and symptoms of infection who meet CDC criteria for SARS-CoV-2 testing. | EUA | https://www.maccura.com/en/ | https://www.maccura.com/resources/98099844450616688.pdf | Swab | PCR | No | Applied Biosystems 7500 | 120 | High | Not specified | 64
| Maccura Biotechnology (USA) LLC | Maccura Biotechnology (USA) LLC | SARS-CoV-2 Fluorescent PCR Kit | BUSGN7103109 | SARS-CoV-2 Fluorescent PCR Kit is a real-time RT-PCR test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in orpharyngeal swab and sputum specimens from individuals with signs and symptoms of infection who meet CDC criteria for SARS-CoV-2 testing. | EUA | https://www.maccura.com/en/ | https://www.maccura.com/resources/98099844450616688.pdf | Swab | PCR | No | Applied Biosystems 7500 | 120 | High | Not specified | 96
| Mesa Biotech | Mesa Biotech (Sekisui Diagnostics?- tlm) | Accula SARS-CoV-2 Test |  | SARS-CoV-2 in nasopharyngeal swab samples. Designed for use on the MAGPIX® System, the NxTAG CoV panel | EUA | https://www.mesabiotech.com/coronavirus | https://static1.squarespace.com/static/5ca44a0a7eb88c46af449a53/t/5e8d218517114d6ea7f60dc2/1586307464282/60061+Rev+1+Accula+SARS-CoV-2+IFU.pdf | Swab | PCR | Yes | Accula Dock, Silaris Dock | 30 | Moderate, High | Not specified | Not specified
| NeuMoDx Molecular, Inc. | NeuMoDx Molecular, Inc. | NeuMoDx SARS-CoV-2 Assay | Not specified | Not specified | EUA | https://www.neumodx.com/sars-cov-2/ | Not specified | Swab | PCR | No | NeuMoDx 288 or NeuMoDx 96 | 60 | Moderate, High | Not specified | Not specified
| OSANG Healthcare | OSANG Healthcare | GeneFinder COVID-19 Plus RealAmp Kit | IFMR-45 | The GeneFinder COVID-19 Plus RealAmp Kit is a real-time reverse transcription-PCR assay for the qualitative detection of SARS-CoV-2 nucleic acids in nasopharyngeal, oropharyngeal, nasal, and mid-turbinate nasal swab specimens, bronchoalveolar lavage fluid (BAL), and sputum from individuals who are suspected of COVID-19 by their healthcare provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 USC §263a, to perform high complexity tests. | EUA | http://www.osanghc.com/en/products_en/molecular-diagnosis/# | https://www.fda.gov/media/137116/download | Swab | PCR | No | Applied Biosystems 7500/7500FAST, Bio-Rad CFX96 | 120 | High | Not specified | 100
| PerkinElmer, Inc. | PerkinElmer, Inc. | PerkinElmer New Coronavirus Nucleic Acid Detection Kit | Not specified | detects three separate gene targets to ensure accurate, reliable results. | EUA | https://perkinelmer-appliedgenomics.com/home/products/new-coronavirus-2019-ncov-nucleic-acid-detection-kit/ | https://www.fda.gov/media/136410/download | Swab | PCR | No | Applied Biosystems 7500 | Not specified | High | Not specified | 48
| Primerdesign Ltd. | Primerdesign Ltd. | Primerdesign Ltd COVID-19 genesig Real-Time PCR assay | Not specified | The genesig Real-Time PCR Coronavirus (COVID-19) CE IVD assay is an in vitro diagnostic test based on Real-Time PCR technology, developed for specific detection of SARS-CoV-2 viral RNA. The probe system is based on the standard hydrolysis probe system known as TaqMan® Technology. The COVID- 19 specific probe is labelled with the FAM fluorophore and the internal control is labelled with the HEX fluorophore. | EUA | https://www.genesig.com/home | https://www.genesig.com/assets/files/Path_Coronavirus_COVID_19_CE_IVD_IFU_Issue_10.pdf | Swab | PCR | No | Applied Biosystems 7500 | 120 | Not specified | Not specified | Not specified
| QIAGEN GmbH | QIAGEN GmbH | QIAstat-Dx Respiratory SARS-CoV-2 Panel | 691223 | QIAstat-Dx Respiratory SARS-CoV-2 Panel is a multiplexed nucleic acid test intended for use with QIAstat-Dx system for the simultaneous qualitative detection and identification of multiple respiratory viral and bacterial nucleic acids in nasopharyngeal swabs (NPS) eluted in universal transport media (UTM) obtained from individuals with signs and symptoms of infection or who are suspected of COVID-19 by their healthcare provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests. | EUA | https://qiastat-dx.com/na/qiastat-dx-sars-cov-2-eua-2/ | https://www.qiagen.com/us/resources/resourcedetail?id=250fafc2-3e66-4400-b5cb-ca69ba8607ab&lang=en | Swab | PCR | No | QIAstat-Dx Analyzer | 60 | High | Not specified | 6
| Quest Diagnostics Infectious Disease, Inc | Quest Diagnostics Infectious Disease, Inc | Quest SARS-CoV-2 rRT-PCR test | Not specified | The Quest Diagnostics SARS-CoV-2 RNA, Qualitative Real-Time RT-PCR (“Quest SARS-CoV-2 rRT-PCR”) is a real-time RT-PCR test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in upper and lower respiratory specimens (such as nasopharyngeal or oropharyngeal swabs, sputum, tracheal aspirates, and bronchoalveolar lavage) collected from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to Quest Diagnostic Laboratories or other laboratories designated by Quest Diagnostics that are also certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. § 263a, to perform high complexity tests. | EUA | https://testdirectory.questdiagnostics.com/test/test-detail/39433/sars-cov-2-rna-qualitative-real-time-rt-pcr?q=39433&cc=MASTER | https://www.fda.gov/media/136231/download | Swab | PCR | No | Applied Biosystems 7500 | Not specified | High | Not specified | Not specified
| Quidel Corporation | Quidel Corporation | Lyra SARS-CoV-2 Assay | M120 | Lyra SARS-CoV-2 Assay (includes controls) | EUA | https://www.quidel.com/molecular-diagnostics/lyra-sars-cov-2-assay | Not specified | Swab | PCR | No | bioMerieux NucliSens easyMAG system and Applied Biosystems 7500 | 75 | High | Not specified | 96
| Quidel Corporation | Quidel Corporation | Lyra SARS-CoV-2 Assay | M935 | Lyra SARS-CoV-2 Assay and Transport Media with Flocked Swab Set | EUA | https://www.quidel.com/molecular-diagnostics/lyra-sars-cov-2-assay | Not specified | Swab | PCR | No | bioMerieux NucliSens easyMAG system and Applied Biosystems 7500 | 75 | High | Not specified | 96
| Quidel Corporation | Quidel Corporation | Lyra SARS-CoV-2 Assay | M936 | Lyra SARS-CoV-2 Assay and Transport Media with Mini-tip Flocked Swab Set | EUA | https://www.quidel.com/molecular-diagnostics/lyra-sars-cov-2-assay | Not specified | Swab | PCR | No | bioMerieux NucliSens easyMAG system and Applied Biosystems 7500 | 75 | High | Not specified | 96
| Roche Molecular Systems, Inc. | Roche Molecular Systems, Inc. | Cobas | 09175431190 | SARS-COV-2 KIT COBAS 6800/8800 SARS-COV-2 192T | EUA | https://diagnostics.roche.com/global/en/products/params/cobas-sars-cov-2-test.html | https://www.fda.gov/media/136049/download | Swab | PCR | No | Cobas 6800/8800 | Not specified | High | Copan UTM-RT, BD UVT | 192
| Roche Molecular Systems, Inc. | Roche Molecular Systems, Inc. | Cobas | 09175440190 | SARS-CoV-2 Positive Control Kit | EUA | https://diagnostics.roche.com/global/en/products/params/cobas-sars-cov-2-test.html | https://www.fda.gov/media/136049/download | Swab | PCR | No | Cobas 6800/8800 | Not specified | High | Copan UTM-RT, BD UVT | 16
| Roche Molecular Systems, Inc. | Roche Molecular Systems, Inc. | Cobas | 07002238190 | 6800/8800 Buffer Negative Control Kit | EUA | https://diagnostics.roche.com/global/en/products/params/cobas-sars-cov-2-test.html | https://www.fda.gov/media/136049/download | Swab | PCR | No | Cobas 6800/8800 | Not specified | High | Copan UTM-RT, BD UVT | 16
| ScienCell Research Laboratories |  | ScienCell SARS-CoV-2 Coronavirus Real-time RT-PCR (RT-qPCR) Detection Kit | RX7038 | ScienCell™ SARS-CoV-2 Coronavirus Real-time RT-PCR (RT-qPCR) Detection Kit (catalog #RX7038, 100 specimens) is a real-time RT-PCR test intended for the qualitative detection of nucleic acids from the SARS-CoV-2 in nasal, nasopharyngeal, oropharyngeal swab specimens, and bronchoalveolar lavage from individuals suspected of COVID-19 by their healthcare provider. | EUA | https://www.sciencellonline.com/sars-cov-2-coronavirus-detection-kit | https://www.sciencellonline.com/media/pdf/SCell-SARS-ifu.pdf | Swab | PCR | No | Roche LightCycler® 96 | Not specified | High | Not specified | 100
| Thermo Fisher Scientific, Inc | Thermo Fisher Scientific, Inc | TaqPath COVID-19 Combo Kit | A47814 | Applied Biosystems TaqPath COVID-19 Combo Kit | EUA | https://www.thermofisher.com/us/en/home/clinical/clinical-genomics/pathogen-detection-solutions/coronavirus-2019-ncov/genetic-analysis/taqpath-rt-pcr-covid-19-kit.html | https://assets.thermofisher.com/TFS-Assets/LSG/manuals/MAN0019181_TaqPath_COVID-19_IFU_EUA.pdf | Swab | PCR | No | Applied Biosystems 7500 | 180 | High | Not specified | 1000
| Trax Management Services Inc. | Trax Management Services Inc. | PhoenixDx 2019-CoV | PCCSKU15259 | PhoenixDx® 2019-nCoV is a real-time Qualitative RT-PCR-based test intended for the qualitative detection of nucleic acid from SARS-CoV-2 in nasal, nasopharyngeal and oropharyngeal swabs and BAL specimens from individuals suspected of COVID-19 by their healthcare provider. Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, to perform high complexity tests. | EUA | Not specified | https://www.fda.gov/media/137153/download | Swab | PCR | No | BIO-RAD CFX96-IVD platform | Not specified | High | Not specified | 50
| Wadsworth Center, New York State Department of Public Health's (CDC) | Wadsworth Center, New York State Department of Public Health's (CDC) | CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel | Not specified | Not specified | EUA | Not specified | Not specified | Swab | PCR | No | Not specified | Not specified | High | Not specified | Not specified
No EUA | RayBioTech | RayBioTech | Coronavirus (SARS-CoV-2) Real Time RT-PCR Nucleic Acid Detection Kit | PCR-COV | The Coronavirus (SARS-CoV-2) Real Time RT-PCR Nucleic Acid Detection Kit is based on the PCR method which uses a fluorescent probe and a specific primer to detect three specific regions within the novel coronavirus (SARS-CoV-2) nucleocapsid protein N gene. This molecular panel aids in the detection of viral RNA from SARS-CoV-2, the causative agent of COVID-19. The kit includes 3 primer-probe sets corresponding to those used in the CDC 2019-Novel Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic Panel (cat no. 2019-nCoVEUA-01). | No | https://www.raybiotech.com/covid-19-rt-pcr/ | https://www.raybiotech.com/files/manual/PCR/PCR-COV.pdf | Swab | PCR | No | Not specified | Not specified | High | Not specified | 20
| Laboratory Corporation of America (LabCorp) | Laboratory Corporation of America (LabCorp) | COVID-19 Test (At-Home Kit) | Not specified | Not specified | Pending | https://www.pixel.labcorp.com/at-home-test-kits/covid-19-test | Not specified | Swab | PCR | No; At-home test | Not specified | Not specified | Not specified | Not specified | 1
| OraSure Technologies | OraSure Technologies | Not specified | Not specified | Not specified | No | Not specified | Not specified | Not specified | Not specified | No; At-home test | Not specified | Not specified | Not specified | Not specified | Not specified